L2 Controller Board Rev.2, User Manual

| Author | Laboratory |  | Approved by | Laboratory |
| --- | --- | --- | --- | --- |
| Karl-Heinz Sulanke DESY |  |  |  |  |

| List of Abbreviations |  |  |  |
| --- | --- | --- | --- |
| SPI | Serial Peripheral Interface | RW | Read / Write |
| L2CB | L2 Controller Board | RO | Read Only |
| CTDB | Clock & Trigger Distribution Board | RWC | Read / Write Clear |
| L2BP | L2 Crate Backplane | 76h | 76 hexadecimal |
| LVDS | Low Voltage Digital Signalling |  |  |

| History |  |  |
| --- | --- | --- |
| Version | Date | Observation |
| 1 | 18/10/2016 | Draft |
| 2 | 20/03/2017 |  |
| 3 | 29/05/2018 |  |

| Distribution |  |
| --- | --- |

Table of Contents

1	Introduction	3

2	L2CB Interfaces	3

2.1	Ethernet to FPGA	3

3	Slow Control via SPI Bus	4

4	L2CB Register Address Map, Base Address = 000h	6

5

Introduction

This document describes the slow control software interface between the L2CB board and the CTDBs.

# L2CB Interfaces

## Ethernet to FPGA

The microcontroller based unit Stamp9G45 is used as a 100MBit Ethernet to FPGA-bridge. It shares a 16 bit memory bus interface with the FPGA.

Figure 1: Stamp9G45

To communicate with each of the 18 CTDBs a shared SPI bus is implemented. The L2CB is the bus master.

## L2CB to TIB

The L2CBs upper RJ45 connector, named “TIP_clock”,  receives the 10 MHz clock and the 1PPS signal from the TIB. It is assumed, that the leading edge of the 1PPS signal follows the 10MHz leading edge after 7ns.

# Slow Control via SPI Bus

Figure 2: SPI Bus Modes. The used SPI Bus mode is of the type CPOL = 0, CPHA = 1

| L2 Backplane, SPI Bus signals |  |
| --- | --- |
| SPI_SCLK | L2CB output that provides 32 clock cycles at each slow control access. The frequency of the clock it is 6.25 MHZ . The electrical standard for this signal is LVDS. Other frequencies are possible. |
| SPI_SYNCn | L2CB output , low during the slow control access and else high. Electrical standard for this signal is LVDS. |
| SPI_MOSI | Master output / slave input data line. A 32-bit word is transmitted at each slow control access. Electrical standard for this signal is LVDS. The most significant bit is transmitted first. The format of the 32-bit word is shown in table 1. |
| SPI_MISO | Master input / slave output data line. Electrical standard for this signal is LVDS. A 16-bit word is transmitted at each slow control access. The most significant bit is transmitted first. Active at a read access, else tristated. This line provides the information of the addressed CTDB register. |

| Bit range | Usage |
| --- | --- |
| 31 | RD = ‘0’, WR = ‘1’ |
| 30-16 | Address_14..0, register address |
| 15-0 | Data_15..0, data word |

Table 1: Slow control 16 bit data field, Bit-15 (MSB) is sent first

|  | SPI_SYNCn |
| --- | --- |
|  | SPI_MOSI |
|  | SPI_MISO |
|  | SPI_SCLK |
| Figure 3: Example, writing data 4321h to address 4f55h |  |
|  | SPI_SYNCn |
|  | SPI_MOSI |
|  | SPI_MISO |
|  | SPI_SCLK |
| Figure 4: Example, reading data 8765h to address 68aah |  |

# L2CB Register Address Map, Base Address = 000h

| hex addr. | Name | Usage |
| --- | --- | --- |
| 00 | CTRL | Control register |
| 02 | STAT | Status register |
| 04 | SPAD | SPI Address register |
| 06 | SPTX | SPI Tx Data register |
| 08 | SPRX | SPI Rx Data register |
| 0a | TSTMP0 | Time stamp, bits 15..0 |
| 0c | TSTMP1 | Time stamp, bits 31..16 |
| 0e | TSTMP2 | Time stamp, bits 47..32 |
| 10 | TEST | Test register, 16 bit R/W |
| 12 | L1SEL | L1 Trigger channel select, for masking and up delay adjust |
| 14 | L1MSK | L1 Trigger mask for channels of selected L2-crate slot |
| 16 | L1DEL | L1 trigger delay in 37ps steps, 0..5ns |
| ... |  |  |
| ... |  |  |
| fe | FREV | Firmware Revision, bits 15..0 |

The registers can be accessed by the Stamp9G45 module via ethernet or RS232

Register Description

| attribute | explanation |
| --- | --- |
| RW | Read / Write |
| RO | Read Only |
| RWC | Read / Write Clear (writing a ‘1’ clears the bit) |

CTRL @ address 00h, 16 bit RW, power_on_value = 0 (0000h)

| Bit | Attr. | Usage |
| --- | --- | --- |
| 15 | RW | 0 to 1 change latches a 48 bit time stamp counter, clocked by TIB_10MHz |
| 14..0 | RW | reserved |

STAT @ address 02h, 16 bit RO, RWC, power_on_value = 0 (0000h)

| Bit | Attr. | Usage |
| --- | --- | --- |
| 0 | RO | ‘1’= SPI bus cycle busy |
| 1 | RO | ‘1’= L1 delay set busy |
| 15..2 | ‘0’ | Reserved, read value is ‘0’ |

Status bits to poll on the readiness of the operation.

Any write operation to the SPAD register initiates a SPI cycle. For SPI write operations, data have to be written to the SPTX register first !  Before initiating the next SPI cycle check STAT-bit_0 (polling on readiness).

SPAD  @ address 04h, 13 bit RW, power_on_value = 0 (0000h)

| Bit | Attr. | Usage |
| --- | --- | --- |
| 7..0 | RW | CTDB register address |
| 12..8 | RW | L2-Crate, CTDB slot number: 1..9, 13..21 |
| 14..13 | ‘0’ | Reserved, write ‘0’ |
| 15 | RW | ‘1’ = Write cycle, ‘0’ = Read cycle |

SPTX  @ address 06h, 16 bit RW, power_on_value = 0 (0000h)

| Bit | Attr. | Usage |
| --- | --- | --- |
| 15..0 | RW | SPI (CTDB register ) data, to be sent |

SPRX  @ address 08h, 16 bit RO, power_on_value = 0 (0000h)

| Bit | Attr. | Usage |
| --- | --- | --- |
| 15..0 | RO | SPI (CTDB register ) data received |

The 0 to 1 transition of  the CTRL-reg. bit_15 latches the 48 bit counter, driven by a 125MHz clock (8ns). The clock is synchronous to the TIB-10MHz clock.

TSTMP0  @ address 0ah, 16 bit RO, power_on_value = 0 (0000h)

| Bit | Attr. | Usage |
| --- | --- | --- |
| 15..0 | RO | time stamp, bits 15..0 , multiples of 8ns |

TSTMP1  @ address 0ch, 16 bit RO, power_on_value = 0 (0000h)

| Bit | Attr. | Usage |
| --- | --- | --- |
| 15..0 | RO | time stamp, bits 31..16 |

TSTMP2  @ address 0eh, 16 bit RO, power_on_value = 0 (0000h)

| Bit | Attr. | Usage |
| --- | --- | --- |
| 15..0 | RO | time stamp, bits 47..32 |

TEST @ address 10h, 16 bit RW, power_on_value = 0 (0000h)

| Bit | Attr. | Usage |
| --- | --- | --- |
| 15..0 | RW | read / write test register, no additional functionality |

L1SEL @ address 12h, 16 bit RW, power_on_value = 0 (0000h)

| Bit | Attr. | Usage |
| --- | --- | --- |
| 3..0 | RW | CTDB channel number of selected slot, valid are: 1..15 |
| 8..4 | RW | L2-crate slot number, valid are: 1..9, 13..21 |
| 15..9 | RO | unused, read value is ‘0’ |

After power on, all L1s are blocked. Write a ‘1’ to the selected L1MSK bit for enabling the dedicated L1 signal.

L1MSK @ address 14h, 16 bit RW, power_on_value = 0 (0000h)

| Bit | Attr. | Usage |
| --- | --- | --- |
| 0 | RO | unused, read value is ‘0’ |
| 15..1 | RW | trigger mask for the selected (L1SEL reg.) slot, each channel has its own select bit, valid are 0000h...fffeh, ‘0’ means L1-trigger of the selected channel is not used |

Before changing the L1 delay of a particular channel, check  the STAT-bit_1 (polling on readiness). The delay for other channels can be set immediately.  Note, the readiness for a particular channel requires the occurrence of several L1 signals.

L1DEL @ address 16h, 8 bit RW, power_on_value = 0 (0000h)

| Bit | Attr. | Usage |
| --- | --- | --- |
| 7..0 | RW | L1 delay of selected ( by L1SEL reg.) trigger channel, usable range 0..128 (x 37ps) |
| 15..8 | RO | 00h |

FREV @ address fe, 16 bit RO

| Bit | Attr. | Usage |
| --- | --- | --- |
| 15..0 | RO | read back of the L2CB firmware revision |