## Slide 1: In Water Operations in the 2025 / 26 Season, SpoolHub Cabling

## Slide 2: SpoolHub-Powering by Batteries
  Fully decouple the SpoolHub from the EMI generating spool-drive
  +/- 48 battery power over the slip ring
  miniPC powered by the mini-FBAP
Temp.
control
Charger
Heater
2x Battery
Slip ring
temperature-isolated box
110 VAC
SpooHub, by J. Kelley
12V

## Slide 3: SpoolHub
  miniPC powered by the mini-FBAP
Quad #1
12V
96V
30W
4W+
ethernet
USB
Enclosure 400x230x111 mm
Paro
RS 485

## Slide 4: Spoolhub, Cables
Item | Conn. #1 | Conn. #2 | Cable type | Length | Comment
Paro to RevPI 4 | PT07A-14-5S | Cable end sleeve | 969M101-22-2TPS-025 | 0.5m | Cable type / length is variable
Mini-FBAP to quad-conn. | DSUB9-male, Norcomp 772-E09-103R001 or Phoenix VS-09-ST-DSUB-EG or TE 5-747904-2 (or -5) | PT07A-14-12S | 969M101-22-2TPS-025 | 0.5m | Cable type / length is variable, DSUB9 housing: Molex 173111-0496, -35°C to 125°C
96V power input to FBAP-mini | PT07A-14-5P | 1-178128-3 | 969M101-14-3-050 | 0.5m | Crimp pins 1-917511-2 14AWG, cable type / length is variable
Ethernet connection | Weidmuller 1018790000 incl. protection cap | RevPI 4 built-in | Telegaertner TG 100007966 | 0.5m | Cat6A, -40°C to +75°C, see Weidmuller spec., reduce (milling) the wall thickness down to 3mm
USB-A to USB-B cable | USB-A male | USB-B male, right angle, down (!) pointing | DeLock 84809 | 0,5m | Temperature range?
RevPI 4 power cable | Molex 22-01-3027 Crimp pins 08550102 | Cable end sleeve | Silicon, stranded, twisted, flex 22AWG, | 0.5m | FBAP-mini, J19, Molex 22-29-2021
WLAN-Antenna TANGO23/0.5M/SMAM/S/RP/19 | n.a. | SMA plug (is part of the antenna) | 50ohm coaxial | 0.5m | Antenna Through Hole Mount, WiFi, WLAN, 2.4 GHz to 5.9 GHz, 5 dBi,

## Slide 5: SpoolHub, Connector Arrangement #1

## Slide 6: SpoolHub, Connector Arrangement #2

## Slide 7: Spoolhub, Status
  Available
    enclosure
    FBAP-mini (incl. new 30W DC/DC)
    FDOR cards
    MS connectors (?)
    Panel mount ethernet connector (?)
    Industry PC „RevPI connect 4“, WLAN | 8GB | 32GB
    USB cabling
  Ordered
    Mini-crate parts by nVent | Schroff
  Still open
    Mini-crate assembly
    Cable (-conn.) details
    WLAN antenna
    Enclosure, connectors holes …

## Slide 8: Battery Box, Status
  Available
    Zarges aluminum box, 600x400x340 mm
    2x LiFePO4 battery, 48V 50Ah, with built-in BMS (Battery Management System)
    1x battery charger
    2x heating mats, 50W
    Charger on/off switch
  Open items
    Thermal insulation
    second battery charger
    Internal cabling
    External connectors
    External cabling