## Slide 1: mDOT_chan0b Test
Kalle Sulanke
May 2020

## Slide 2: Introduction
Available now: 4 x mDOT_chan0b + 5 x mDOT_chan0, assembled by Mandy Hemp
chan0b almost identical with chan0, except
  Optional, extra DAC channel could be used to tune the preamps working point
    Not used presently
  additional input and improved ADC protection
Present tests with bigger time constant for the pulse shaper
  91ohm x 100pF instead of 49.9ohm x 100pF (like at chan0 and chan2)

## Slide 3: LTSpice Schematic
time constant

## Slide 4: SPE Signal, LTSpice Simulation, 1SPE
29
ADC bins

## Slide 5: SPE Signal, LTSpice Simulation, 1-2 SPE
ADC_in
1SPE
2SPE
58
ADC bins

## Slide 6: SPE Signal, LTSpice Simulation, 1-80 SPE, Linearity
ADC_in
2396 ADC bins
Linearity, zoomed view

## Slide 7: SPE Signal, LTSpice Simulation, 2x 1SPE, 15ns

## Slide 8: SPE Signal, LTSpice Simulation, 2x 1SPE, 10ns

## Slide 9: Typical noise, with open input
Rin=910R, Rn = 180R, Cn=610p, LP = 91R x 100p, gain_2 ~ 5, DAC0_e = 37000, DAC1_a = 37200

## Slide 10: 1/4 SPE input (2mV pk)
6
100% trigger rate

## Slide 11: 1SPE input (8mV pk)
24

## Slide 12: PMT signal, HV = 800V, discr_thr = 1SPE
42
PMT-Testbox used
  R5900-3-M4
  Typ. HV: 800 +/-50V

## Slide 13: PMT signal, HV = 1000V (!), discr_thr = very high
PMT-Testbox used
  R5900-3-M4
  Typ. HV: 800 +/-50V
DAC1_a = 55000
3300

## Slide 14: Double Pulse, HV = 800V, discr_thr = 0.1 SPE
PMT-Testbox used

## Slide 15: Double Pulse, HV = 900V, discr_thr = 0.25 SPE
PMT-Testbox used